## Traceability Matrix
Requirements Traceability Matrix
Source, one column per hop, then Coverage and Result. Coverage and Result are calculated; do not type in them. See the How to use sheet.
Source | Requirements (SRS) | Test Protocol (TP) | Test Report (TR) | Coverage | Result
UN-001 | SRS-001 | TP-101 | TR-101
UN-002 | SRS-003 | TP-104
UN-003 | SRS-008 | TP-106 | TR-106
UN-004 | SRS-004 | TP-108 | TR-108
UN-005 | SRS-011 | TP-112 | TR-112
UN-006
TP-120
TP-120 has an empty Source on purpose: nothing reaches it. Traceable lists rows like this at the bottom of the matrix automatically.

## User Needs (UN)
User Needs
One document, one ID grid. These rows are the Source column of the matrix.
Trace ID | User need | Owner | Notes
UN-001 | The device continuously monitors heart rate over a 72 hour wear cycle | A. Demir
UN-002 | The device alerts the user within 5 seconds of a detected arrhythmia | A. Demir
UN-003 | The companion app displays the most recent 24 hours of rhythm data | A. Demir
UN-004 | The device runs for seven days between charges | A. Demir
UN-005 | The wearer can export a report for their clinician | M. Chen
UN-006 | The device warns the wearer before the electrode patch expires | M. Chen | No requirement written yet

## Requirements (SRS)
Software Requirements Specification
Verification method, owner and notes live HERE, on the requirement's own row, not in the matrix.
Trace ID | Requirement | Verification method | Owner | Notes
SRS-001 | The sensor sustains a 256 Hz sampling rate across the recording window | Test | R. Farhadi
SRS-003 | Arrhythmia detection latency is below 5 seconds at the 95th percentile | Test | R. Farhadi
SRS-004 | Battery endurance is seven days under nominal load | Test | R. Farhadi
SRS-008 | The companion app refreshes at an interval no greater than 2 seconds | Demonstration | M. Chen
SRS-011 | A clinician report exports as a PDF within 30 seconds | Test | M. Chen

## Test Report (TR)
Test Report
Result is PASS, FAIL, or left blank when the test has not been run. The matrix reads this column.
Trace ID | Test case | Result | Actual output | Deviation
TR-101 | TP-101 | PASS | 256 Hz held for 72 h, no dropped samples
TR-106 | TP-106 | PASS | Refresh measured at 1.4 s median
TR-108 | TP-108 | FAIL | 6 d 4 h endurance at nominal load | Below the 7 day requirement; retest after firmware rev C
TR-112 | TP-112 |  |  | Executed but result not recorded

## How to use
How to fill in this matrix
Seven steps. The last one is the one that fails.
The matrix has six columns, and you only type into four of them.
Source, one column per hop, then Coverage and Result. Coverage and Result are calculated here and derived in Traceable, so they are never typed by hand in either place.
1. Give every row a Trace ID, in its own document.
User needs on the User Needs sheet, requirements on the Requirements sheet, results on the Test Report sheet. Each sheet stands for one document. Never reuse an ID and never renumber: the ID is what both ends of a trace point at.
2. Put the requirement's own attributes on its own row, not in the matrix.
Verification method, owner, notes and the requirement text belong on the Requirements sheet. In Traceable these are columns on the requirement's row inside its document, and the matrix shows them inside the linked chip. The matrix itself carries IDs, coverage and result, nothing more.
3. Fill the matrix by following links, one hop at a time.
A row starts at a source ID and steps through each hop: which requirement realises this need, which protocol verifies that requirement, which report records the run. Put the Trace IDs reached in the hop cell, separated by "; " when a hop reaches more than one.
4. Read Coverage rather than setting it.
Coverage is Pass when the last hop is populated and Gap otherwise. A row that stops halfway is a Gap, whatever intermediate hops it filled in. That is exactly the rule Traceable applies.
5. Read Result rather than setting it.
Result rolls up the Test Report: any FAIL makes the row Fail, otherwise any PASS makes it Pass, otherwise a result that exists but was never recorded reads Not run. A row with no result at all stays blank.
6. Add a row for anything nothing reaches.
TP-120 on the matrix sheet has an empty Source. It is a test case no requirement points at, which is a coverage gap in the other direction. Traceable finds these for you and lists them at the bottom of the matrix; in a spreadsheet you have to notice them.
7. Re-check the whole matrix after every change.
This is the step that fails, and it fails silently. Nothing here knows that a protocol was retired or a requirement reworded, so a stale row keeps reading as covered until a person checks it by hand.
What this file cannot do, and Traceable does
• One hop cell, one lookup. Result here is calculated from a single Test Report ID per row. Traceable rolls the result up across the whole chain, however many links each hop reaches, which no formula in this file can do.
• Links live in two places at once. Writing SRS-003 in a hop cell does not tell the Requirements sheet anything. In Traceable a link is one object with two ends, so it cannot be right in one direction and wrong in the other.
• Nothing marks a broken link. If TP-104 is retired, this file keeps showing it. Traceable tracks each link's health and flags it instead of leaving the row reading as covered.
• The hop columns are fixed once you save. In Traceable they are configuration: choose the source documents, add or reorder hop columns, save the view, and every row rebuilds from the current links.
Traceable builds this same matrix from live trace links: pick the source documents, configure the hop columns, and every row, its coverage and its result are derived from the links as they stand. Its CSV export has the same columns as the matrix sheet in this file. https://traceabledocs.com/requirements-traceability-matrix

## Lists
Verification method | Test Result | Coverage (calculated) | Result (calculated)
Test | PASS | Pass | Pass
Inspection | FAIL | Gap | Fail
Analysis |  |  | Not run
Demonstration